Zeiss Observer Z1 with apotome – Quick Reference Guide for users

For any help, please contact Jiten Manji

Login to computer:

Username: REGMAPOT

Password: 123MICmic

Start-up Procedure

1) Turn on all ‘Power Units’

- Power Supply

- HBO Lamp

- Apotome (optional)

  - Please note if you intend on using the Apotome, this MUST be switch on before using the software

2) Switch on the Microscope by pressing the switch found on the

Camera

left side of the microscope

Focussing Wheel

FRONT

REAR

ON/OFF Switch

3) Open the AxioVision Software on the desktop

4) Choose the appropriate stage insert

Available Microscope Inserts: Image of inserts

- Non-heated multiplate insert

- Non-heated slide insert

5) To load a sample, follow these steps:

- Press ‘Load Position’ from the microscope computer panel

- The objectives will move out of place

- Title the lamp support column backward to allow the stage insert to be correctly placed

- Remove the slide holder  and insert your slide

- Replace the holder back onto the microscope

- Turn the focus wheel and select ‘Set Working Position’ on the microscope control panel

6) Place the stage correctly into the Stage Insert Clips to ensure that the stage is not damaged

RIGHT

LEFT

Objectives

- Live – Opens a window with a live camera feed

- Snap – Captures an image of the live camera feed

- Scale bar – Annotates a calibrated scale bar onto an opened image

- Text – Annotates with text onto an opened image

- Navigator – Defines a viewing region in images and live camera feeds

- Workflow – Opens a defined path of work for achieving specific aims

- Workarea – Opens window for all options for working

- Splitter – Opens a viewing mode where multiple images/series can be viewed simultaneously

- Gallery – Allows selection of which opened/captured image/series to view

- Dapi/GFP/txred – All three will set the microscope to image in the selected channel

- Closed – turns all imaging modes off and closes shutters

- Incubation – Opens settings window for heated elements and atmospheric controls

Control Section

Control Section

Microscope Controls

(these can also be accessed via touchscreen control panel)

Current Objective

Multi-dimensional Acquisition of Images

Control Tabs

Option Selections

Re-use a previously saved experiment

The most common way to acquire a multi-parameter image is to use the Multi-dimensional acquisition tab found within the workarea. These options can be used separately or in conjunction with one another.  I.e. 3 Colours, multiple points and time lapse etc…

Multichannel – multiple coloured fluorophores

Z-Stack – Imaging through the Z plane

Time Lapse – Multiple time points

Position List – Multiple positions

C-Tab for multichannel imaging

The default setting is for all 3 fluorescent channels and a phase contrast ‘channel’. These can be de-selected by right clicking on the separate channel boxes at the top.

Channel Selection

Dye Definition

Exposure Measurement

Hardware Settings

Important: Please double check that the right filter cube has been selected

This shouldn’t normally need to be changed

1) Select desired channels depending on the fluorochromes used in you protocol

- You may also refer to the fluorochrome table on the desk

2) Channels can be de-activated by ‘right-clicking the tab

3) Click on measure and adjust the camera exposure accordingly for each channel separately

Z-tab for imaging through the z plane

You can stipulate the number of stacks desired, or you can allow the system to calculate depending on the thickness of your specimen.

Focussing window

Stack dimension selection options

- Choose Start/stop and open the focussing window next to the ‘Start’ box

- This window provides a live image feed of the sample which can be focussed manually or using the slider to the left of the window.

- Focus to a position above cells

- Click ok in focussing window

- Repeat for ‘Stop’ by focussing beneath sample/cells

T-tab for Time Lapse microscopy

Open the T tab if you desire to collect a time lapse/live video tick the ‘Time Lapse’ Box

Select the appropriate interval and duration/cycle number for time lapse experiment – Remember that if the interval is too short to complete all commands the interval will lengthen regardless.

Select the Interval

Number of cycles

Apotome

In thick specimens, the signal from the focal plane is compromised by the out of focus light coming from planes above and bellows this plane. The result is a blurred, low contrast image. In this condition, is not possible to render a three-dimensional projection of the specimen. Some techniques were developed to remove the

out of focus light, either optically (Confocal) or mathematically (Apotome).

In the Apotome, a grid is inserted in the illumination pathway of a conventional fluorescence microscope. A plane-parallel glass plate is placed in front of the grid and it tilts back and forth projecting the grid in the image in 3 different positions. For each depth, 3 raw images with the projected grid superimposed in 3 different positions are taken. The software uses the projected images of the grid to calculate and removes the out of focus light. The principle is, if the grid is visible that information is in focus. If not, that signal corresponds to out of focus light and it is removed from the final image. The software combines the 3 images into 1 for each stack.

To acquire an image using the Apotome

Apotome Power Supply

- Switch on the Apotome power supply (this should be done with the start-up procedure

2) Gently insert eh Apotome Slider (all the way in) until you hear a beep

3) Open the Apotome window

3a) Click measure and the software will calculate the optimal exposure time

3b) Click refresh, you can also crop your image

4) Move to settings tab

4a) Select Grid Visible

4b) Select Optical Sectioning for Z-stacks

4c) Select Conventional Fluorescence for standard imaging

4d) Ensure Grid ‘VH’ is selected. This can be used with both x20 and x40 objectives only. If you require a lower objective please email Jiten as the system will need to be re-calibrated

Trouble shooting with the Apotome

1) Apotome Filtering

2) Apotome Averaging

1) If your final processed image displays a grid pattern, use the ‘Apotome Filtering’ set the option to weak, medium, strong depending on the intensity of the grid.

2) If you find that there is significant noise in your final processed images, increase the number of averages.

Apotome Image

Widefield Image

POST-ACQUISITION OPTIMISATION

Once data has been collected the image can be enhanced so more detail is visible and also blurry z stacks can be reduced to a single plane image which uses in focus from all of the stack.

1 – Brightness/Contrast Enhancement

One of the first enhancements that is nearly always an improvement to an image is to adjust the brightness and contrast of each channel collected so it is the most clearly visible.

- The acquired image/series should be present as a tab just beneath the main toolbar – select the desired image/data

- Save as a ‘.zvi’ file immediately to prevent any data loss – it is best to export the file as desired for publication after saving this way so that if changes are necessary these can be made without data loss.

Fig 25 – Image Window – 2D View

Psuedo-coloring

Tools

Channel Selection

- Turn off Psuedo-colouring switch

- Press Alt+Enter or right click on image and select ‘Properties’ to open an image properties window

Figure 26 - Image Properties/Display Window

Intensity Histogram

Restore Display properties

Automatic Brightness/

Contrasting controls

Save Display properties

Brightness

Contrast

Gamma Sliders

- Alter brightness and contrasting by either moving display indicator directly, using the sliders or using an automatic selection. Linear is the default and is used to reset any changes. The gamma slider (furthest in from edge) causes a curved display indicator which is useful to reduce background or low level staining without removal.

- Choose and repeat for each channel.

2 – Selecting a Region of Interest (ROI)

It is possible to select a smaller area of your image / sequence / series to crop and enlarge. This is defined with the ROI tool.

This is found in the tools section of an image window (Fig 25)

The ROI boundary is defaulted to encompass the entire image

- Select a smaller area by moving the boundaries to the edges of the area desired

- Select Copy ROI in Edit (main tabs) or press ctrl+shift+c.

- Now simply pressing ctrl+v or Edit → Paste and a new image is produced of just the ROI selected.

This process may be especially useful when a long capture has occurred and the file is very large but only a small area has any valuable information.

3 – Annotations

- Adding a scale bar

  - Open image

  - From menu tabs select Annotations → Scale Bar

  - Draw the length of scale bar desired

  - Open Image Properties (Fig 26)

  - Select second tab

  - Adjust colour/size/font as desired

- Adding a time mark

  - Select Annotations → Frequently used annotations → Time Mark

Time marks annotate the time at which images were collected. An alternative is relative time mark which shows sequence duration.

- Remember to select ‘Burn in annotations’ when exporting

4 – Exporting As Another File Type

Once a file has been captured is always best to save it as a .zvi file but it is not always best to transport such large files and mostly these cannot be used to publish.

To export a file it must be opened and being viewed.

- Open File → Export

- Select a file type to export as

- Select the desired percentage of compression – if none is desired select 0%

- Choose a name and location in relevant areas of the Export window.

- Click on Start

The default location for files to be exported is on the desktop.

Please transport files via the X-Drive or use a memory device.

SHUT DOWN

Before exiting the software, Please SAVE ALL DATA onto the shared X:

Once you have saved and exported your files safely the system needs to be shutdown safely.

1) Please check the outlook calendar by logging on to office365 if there is another user booked after you.

1a) If there is a user within 1hour of you finishing, please leave the HBO lamp on and log off the computer

1b) If you are the last user of the day, please shut down the microscope and all components – See below

2) Turn off all components

a) Microscope

b) HBO Lamp

c) Apotome power supply

d) Shutdown computer and turn off monitor

Important, please note that time on and off should be recorded on the sheet located next to the microscope.