## BioSci NPP & CDB (hidden)
UCL Radiation Audit Action Points 2020 - 2021
Audit # | UCL No. | Action Point | Start | Finish | Assign to | Priority | Completion % | Status | Commenets | Column11
Division of Biosciences, Darwin Building
1 | 1 | The signed RPS appointment letter from the Head of Department. The letter must specifically list the labs or departments that he is the RPS for | 2020-10-23 | 2020-12-31 | all RPS | High | 0% | Not Started
Neuroscience, Physiology and Pharmacology 4th Floor of Medical Sciences Building, Room 340
1 | 2 | The designated sink for radioactive waste disposal needs a catch pot placed underneath to reduce the risk of contamination from leaks. The sink and surrounding area needs to be cleared and cleaned and monitored regularly when work with radioactive materials recommences. Stainless steel will require a deep clean with decon-90 on a regular basis due to the attractive properties between P-32 and stainless steel. This can lead to contamination issues and reasonable dose rates. Items relating to other non-active work should not be left or stored in this area. The tap should have a long hose attached to reduce the possibility of splashes. The area under the sink should be kept clear to reduce the risk of contamination incidents. | 2020-10-23 | 2020-12-31 | RPS and all researchers | High | 0% | Not Started
2 | 3 | There should be a clear sign on the door stating that the lab contains a radiation supervised area, with a trefoil symbol and which radionuclides are used (P-32) and the name of the lab RPS and contact details. | 2020-10-23 | 2020-12-31 | RPS | High | 0% | Not Started
3 | 4 | The EP15 mini-monitor stored in this lab (SN – 046788) was last calibrated on 26/01/2016. Contamination monitors require calibration on an annual basis and so will need to be calibrated prior to recommencing work with radioactive materials. See Appendix 1 for details on arranging calibration. | 2020-10-23 | 2021-03-31 | RPS | High | 0% | Not Started
4 | 5 | A spill kit should be assembled and kept in a clear plastic box with a lid. The inventory of the spill kit should be as listed in Appendix 2. The contents of the spill kit should be replenished after each use and audited regularly. | 2020-10-23 | 2020-12-31 | RPS | High | 0% | Not Started
5 | 6 | It is advised that a smaller solid waste bin is used and should be disposed of via the correct routes every three months at least. The existing bins must be collected ASAP and relabelled according to the current system in place. | 2020-10-23 | 2020-12-31 | RPS | High | 0% | Not Started
6 | 7 | The work area in which radioactive work is taking place should be free from clutter and kept clean. The floors require a deep clean before radioactive work is carried out again in the lab. Floors that are not kept clean will make the task of clearing up contamination from a spill more difficult. | 2020-10-23 | 2020-12-31 | RPS | Medium | 0% | Not Started
7 | 8 | A suitable allocation of maximum radionuclide holdings should be made available in this laboratory. Orders and holdings must not exceed the allocated maximums. | 2020-10-23 | 2021-03-31 | RPS | High | 0% | Not Started
8 | 9 | Surface monitoring and/or swab testing for contamination should be completed after each session of working with radioactive materials. The results of which must be recorded (in cps) and kept on record. The records must be checked each month by the RPS for completeness | 2020-10-23 | no end date | RPS and researchers | Medium | 0% | Not Started | Initially to discount long lived contamination of furniture and equipment (H-3) and then when work with radioactive materials re-commences.
9 | 10 | The lab requires updated local rules and risk assessments. These are to be issued by the RPO and completed jointly between the RPO and the RPS. These documents should be read and signed by all staff working in the supervised area. | 2020-10-23 | 2021-03-31 | UCL RPO & RPS | High | 0% | Not Started
10 | 11 | Local standard operating procedures are to be put in place including - Ordering of stock - Record keeping - Identification of radioactive areas - Radioactive waste stores - Accumulation and disposal of solid waste and sealed sources - Accumulation and disposal of aqueous waste - Notification procedure.Spills and Accidental Releases - Design and planning of radioactive workspaces and storage - Decommissioning and clearance of radiation labs - Permit variations, surrenders and new permit applications. - Best Available Techniques (BAT) - Leak Testing - Training and induction of radiation workers - High Activity Sealed Sources - RPS Audit Procedure - Project Approval | 2020-10-23 | 2021-03-31 | UCL RPO & RPS | Medium | 0% | Not Started | All procedure headings must be available – even though some procedures may not be applicable. In this instance the procedure title must be followed by the words ‘not applicable in this department’.
Cell & Development Biology, 2nd Floor of Anatomy Building, Room 209
1 | 12 | The sign on the entrance to the lab should be updated to give the correct RPS details, RPO details, lab designation (supervised area), radionuclides in use and maximum activities stored within the lab. | 2020-10-23 | 2020-12-31 | RPS | Low | 0% | Not Started
2 | 13 | The designated sink for radioactive waste disposal needs a catch pot underneath to reduce the risk of contamination from leaks. The area under the sink should be kept clear to reduce the risk of contamination incidents. A long hose should be placed on the tap into the sink to reduce the likelihood of contamination arising from splashes. | 2020-10-23 | 2021-03-31 | RPS | Medium | 0% | Not Started
3 | 14 | There are 6 radiation monitors stored within the lab. All of which are due for calibration. Please see Appendix 1 for arranging calibration: Monitor SN: 046384, Calibration Due: 09/02/2012 Monitor SN: 025979, Calibration Due: 18/11/2000 Monitor SN: 028885, Calibration Due: 31/10/2008 Monitor SN: E0002602, Calibration Due: 14/11/2009 Monitor SN: 1332, Calibration Due: 26/03/2007 Monitor SN: 023176, Calibration Due: 26/01/2017 | 2020-10-23 | 2021-03-31 | RPS/RPO | HIGH | 0%
4 | 15 | The area in which radiation work is to be done was found to be cluttered. The designated area and bench should be kept clutter free and clean to minimise the risk of contamination and spills. | 2020-10-23 | 2020-12-31 | RPS | Medium
5 | 16 | The spill kit should be transferred to a clear plastic box for easier access. The inventory of which should be as listed in Appendix 2. The contents of the spill kit should be replenished after each use and audited regularly. | 2020-10-23 | 2020-12-31 | RPS | Low
6 | 17 | Local standard operating procedures are to be put in place including: -Ordering of stock, Identification of radioactive areas - Radioactive waste stores - Accumulation and disposal of solid waste and sealed sources - Accumulation and disposal of aqueous waste - Notification procedure - Spills and Accidental Releases - Design and planning of radioactive workspaces and storage - Decommissioning and clearance of radiation labs - Permit variations, surrenders and new permit applications. - Best Available Techniques (BAT) - Leak Testing - Training and induction of radiation workers - High Activity Sealed Sources - Audit Procedure - Project Approval - RPS Auditing Record keeping | 2020-10-23 | 2021-03-31 | RPS & RPO | Medium | 0%
7 | 18 | The scintillation counter should be labelled on the front surface with a trefoil and details of the sealed source. Instructions relating to the disposal of the equipment should also be made clear as follows “Sealed source inside – do not dispose of this equipment without approval from the UCL RPO”. Records of wipe testing should be requested from the servicing company. | 2020-10-23 | 2021-01-31 | RPS | Medium | 0%
8 | 19 | The freezer must be labelled with a trefoil symbol and the possible contents should be listed on the door. The drawer that the radionuclides are kept in must also be labelled. The freezer is not lockable. A lock must be found for this freezer or else another freezer must be found from another location that already has a lock on it. | 2020-10-23 | 31/12/220 | RPS | High | 0%
9 | 20 | Whilst the lab is not being used for hot work the stock that is currently stored on the bench top and shelving must be kept somewhere more secure and away from where people are currently working. Security of these sources is a big issue and the RPS should actively audit these sources to confirm that they are still present. It is not acceptable to keep radioactive sources on a shelf or windowsill even with shielding. | 2020-10-23 | 2020-12-10 | RPS | High | ?
10 | 21 | There are currently no immediate plans to use any of the stock of radioactive material currently stored in the lab. It was agreed that all stock should be disposed of via the correct routes and fresh stock can be ordered when the lab has been cleared and tidied and the correct new procedures and documentation are in place. Please contact the UCL RPO with any queries regarding how best to dispose of stock. | 2020-10-23 | 31/03/20201 | RPS & RPO | Medium | ?
11 | 22 | Surface monitoring and/or swab testing for contamination should be completed after each session of working with radioactive materials. The results of which must be recorded (in cps) and kept on record. The records must be checked each month by the RPS for completeness. Initially to discount long lived contamination of furniture and equipment (H-3) and then when work with radioactive materials re-commences. | 2020-10-23 | 2021-03-31 | RPS & all users at the time of leaving the lab | High | ?
Neuroscience, Physiology and Pharmacology (NPP) 3rd Floor, Andrew Huxley Building, Room 304
1 | 23 | The sign on the entrance to the lab should be updated to give the correct RPS details, lab designation (supervised area), radionuclides in use and the maximum activities that can be held within the lab. | 2020-10-23 | 2020-12-31 | RPS | Medium | ?
2 | 24 | The designated sink for radioactive waste disposal needs a catch pot underneath to reduce the risk of contamination from leaks. The area under the sink should be kept clear to reduce contamination risk. | 2020-10-23 | 2020-12-31 | RPS | Medium | 0% | On going?
3 | 25 | There are 2 radiation monitors stored within the lab. Both of which are due for calibration. Please see Appendix 1 for arranging calibration: Monitor SN: 023491, Calibration Due: 27/01/2016 Monitor SN: 047071, Calibration Due: 01/10/2007 | 2020-10-23 | 2021-03-31 | RPS | High | 0%
4 | 26 | The area in which radiation work is to be done was found to be cluttered. The lab bench should be kept clutter free and tidy to minimise the risk of contamination | 23/10/202 | 2020-12-31 | RPS | Medium | 20%
5 | 27 | The spill kit should be transferred to a clear plastic box with a lid for easier access. The inventory of which should be as listed in Appendix 2. The contents of the spill kit should be replenished after each use and audited regularly. | 2020-10-23 | 2021-03-31 | RPS | Low | 0%
6 | 28 | Local standard operating procedures are to be put in place including procedures for: - Ordering of stock - Record keeping - Identification of radioactive areas - Radioactive waste stores - Accumulation and disposal of solid waste and sealed sources - Accumulation and disposal of aqueous waste - Notification procedure - Spills and Accidental Releases- Design and planning of radioactive workspaces and storage - Decommissioning and clearance of radiation labs - Permit variations, surrenders and new permit applications. - Best Available Techniques (BAT) - Leak Testing - Training and induction of radiation workers - High Activity Sealed Sources - Audit Procedure - Project Approval | 2020-10-23 | 2021-03-31 | RPS & RPO | Medium
7 | 29 | The scintillation counter should be labelled on the front with a trefoil and labelled as containing a Cs-137 sealed source. Wording should include instructions to notify the UCL RPO if the equipment is being moved or disposed. Records of wipe | 2020-10-23 | 2021-01-31 | RPS | High
8 | 30 | There are currently no immediate plans to use any of the stock of radioactive material currently stored in the lab. It was agreed that all stock should be disposed of via the correct routes and fresh stock can be ordered when the desire to restart radioactive work returns. Please contact the UCL RPO with any queries regarding how best to dispose of stock. | 2020-10-23 | 2021-03-31 | RPS & RPO | Medium

## Audits
UCL Action Plan - (Audit)
Audit Ref | UCL No. | Action Point | Start date | End date | Assign to | Priority | Completion % | Status | Comments | Column11
1 | 1 | The signed RPS appointment letter from the Head of Department. The letter must specifically list the labs or departments that he is the RPS for | 2020-10-23 | 2022-12-30 | all RPS | High | 0% | Not Started
2 | 2 | The designated sink for radioactive waste disposal needs a catch pot underneath to reduce the risk of contamination from leaks. The area under the sink should be kept clear to reduce contamination risk. | 2020-10-23 | 2022-08-31 | RPS | Medium | 0% | On going?
4 | 3 | The area in which radiation work is to be done was found to be cluttered. The lab bench should be kept clutter free and tidy to minimise the risk of contamination | 23/10/202 | 2020-12-31 | RPS | Medium | 20% | Incomplete
6 | 4 | Local standard operating procedures (SOP) are to be put in place including procedures for: - Ordering of stock - Record keeping - Identification of radioactive areas - Radioactive waste stores - Accumulation and disposal of solid waste and sealed sources - Accumulation and disposal of aqueous waste - Notification procedure - Spills and Accidental Releases- Design and planning of radioactive workspaces and storage - Decommissioning and clearance of radiation labs - Permit variations, surrenders and new permit applications. - Best Available Techniques (BAT) - Leak Testing - Training and induction of radiation workers - High Activity Sealed Sources - Audit Procedure - Project Approval | 2020-10-23 | 2021-03-31 | RPS & RPO | Low | 40% | Incomplete
Neuroscience, Physiology and Pharmacology (NPP) 3rd Floor, Andrew Huxley Building, Room 304

## Incidences
UCL Action Plan - (Incidences)
Audit ref # | UCL No. | Action Point | Start Date | End Date | Assign to | Priority | Completion % | Status | Comments | Column11

## Inspections
UCL Action Plan - (Inspections)
Audit # | UCL No. | Action Point | Start Date | End Date | Assign to | Priority | Completion % | Status | Comments | Column11

## Summary
SITE | Total # AP | Completed | Ongoing | Not started | High | Medium | Low
1 | 18 | 2 | 3 | 13 | 12 | 6 | 0
2 | 30 | 0 | 3 | 27 | 14 | 13 | 3
| 1 | 2
Completed | 2 | 0
Ongoing | 3 | 3
Not started | 13 | 27